## Slide 1
NeuroSight AI
Longitudinal MRI dementia analysis on OASIS-2
NeuroSight-E5 · 5-seed ensemble · model card 2026-09-17 · report 2026-10-21
Research use only — not a medical device
All numbers from committed run artifacts (reports/) — generated by scripts/build_reports.py

## Slide 2
Problem and approach
Problem
• Binary dementia-status classification per visit (non-demented vs demented) from longitudinal MRI
• OASIS-2: 150 subjects, 373 sessions; labels derived from CDR and verified against the research group field
• Small data regime: honest performance is limited by variance, not by model capacity
• Trust requires calibrated confidence, interpretability, and disclosed failures
Approach
• Per-slice CNN with spatial and channel attention, BiLSTM over visits, additive temporal attention (7_dual_attn)
• Subject-level 70/15/15 split; 5-fold subject-level CV is the primary evidence; the held-out test set was touched exactly once
• mri_only headline mode masks clinical metadata (MMSE and CDR define the label)
• Temperature scaling, 5-seed ensemble, and conformal prediction sets evaluated with pre-registered targets

## Slide 3
Dataset and methodology
• 150 subjects aged 60-96, 373 MRI sessions (Analyze 7.5)
• Audit: 373/373 CSV rows matched to scans on disk; no missing or empty sessions
• Subject groups: 72 nondemented, 64 demented, 14 converted
• Labels: binary per visit from CDR (0 vs above 0), group verified; converted subjects split into pre- and post-conversion sequences
• Subject-level 70/15/15 split + 5-fold subject-level CV (seed 42); leakage guards automated in the test suite
• Preprocessing: z-score, 32 middle axial slices, 128x128
Property | Value
Subjects | 150 (aged 60-96)
MRI sessions | 373 (audit matched 373/373)
Subject groups | 72 / 64 / 14 (nondemented / demented / converted)
Labels | Binary per visit from CDR (0 vs above 0), group verified
Splits | Subject-level 70/15/15 + 5-fold subject-level CV (seed 42)

## Slide 4
Architecture: 7_dual_attn (champion)
Input: T visits x 32 axial slices x 128x128
Per-slice CNN: Conv 1-8-16-32 + BatchNorm + ReLU
CBAM spatial attention + SE channel attention
Slice mean-pool to a 64-d embedding per visit
BiLSTM over visits (hidden 32)
Additive temporal attention (softmax visit weights)
FC head: 2 logits per visit
Deployment: neurosight-e5 (5-seed ensemble)
Mean softmax over members: s42, s43, s44, s45, s46
Temperature fitted on the selection set: T = 0.657

## Slide 5
Honest results: held-out test (n = 22)
Metric | NeuroSight-E5 | Champion v1 *
Accuracy | 59.1% [95% CI 36.4, 77.3] | 59.1%
AUC-ROC | 0.777 | 0.785
F1 (demented) | 0.571 | 0.571
Sensitivity | 0.545 | 0.545
Specificity | 0.636 | 0.636
ECE (temperature-scaled) | 0.227 | 0.304
Fitted temperature | 0.657 | 0.375
* v1-lineage checkpoints predate the training-bug fixes (see disclosures). McNemar chi-squared 0.0, p = 1.0: identical decisions on all 22 test cases.
The accuracy CI covers a 40.9-point range; CV (mean AUC 0.744 ± 0.096) is the primary evidence.

## Slide 6
Cross-validation leaderboard (5-fold, subject-level)
Model | Mean CV AUC | Std
7. DualAttn (champion) | 0.744 | 0.096
10. 3D-CNN+LSTM | 0.732 | 0.098
8. Transformer | 0.727 | 0.075
5. +TemporalAttn | 0.725 | 0.141
11. ResNet+LSTM | 0.710 | 0.172
1. CNN-avg | 0.689 | 0.120
6. +SE | 0.689 | 0.151
9. 3D-CNN | 0.688 | 0.077
4. +SpatialAttn | 0.688 | 0.134
3. CNN+LSTM | 0.685 | 0.118
2. CNN-last | 0.665 | 0.180
• Champion 7_dual_attn carries forward: 0.744 ± 0.096 mean AUC
• Fold-to-fold std spans 0.075 to 0.180: most between-variant gaps are within noise
• All variants are v1 lineage (pre-fix); retained as baselines

## Slide 7
Excellence phase: same-fold systems and negative results
System | Description | Mean AUC | Std | dAUC vs S0
S0 | Champion v1, single model | 0.7444 | 0.0965 | 0.0000
S1a | Same-fold CV, champion checkpoints | 0.7444 | 0.0965 | 0.0000
S1b | Same-fold CV, transformer checkpoints | 0.7270 | 0.0750 | -0.0174
S2 | Cross-variant ensemble (champion + transformer) | 0.7381 | 0.0819 | -0.0063
S3 | S2 plus test-time augmentation | 0.7311 | 0.0873 | -0.0132
• Cross-variant ensembling (S2): 0.006 AUC vs S0 — too correlated to help; dropped
• Test-time augmentation (S3): 0.013 AUC vs S0 — no new information; dropped
• Hyperparameter search (8 configs, 3-fold proxy): top config = deployed config (0.0003 lr, 0.5 dropout, hidden 32; 0.714 ± 0.124) — no change made
• Selection on CV only; the held-out test set was untouched until the bake-off

## Slide 8
Calibration and conformal prediction
Calibration
• Test ECE: 0.304 (v1 lineage) to 0.227 (E5) with temperature T = 0.657 fitted on the selection set
• Raw ensemble without scaling: 0.273
• Above the pre-registered 0.10 target: probabilities are rankings, not calibrated percentages
Conformal sets
• Split-conformal at a 90% coverage target measured 54.5% coverage (mean set size 0.864)
• Empty sets occur at the available calibration size
• Not deployed — recorded as an honest failure in the model card and reports/excellence/

## Slide 9
Disclosures: read before trusting any number
• Cache bug skipped augmentation in the v1 ablation (fixed); v1-lineage baselines inherit it
• Broken gradient accumulation in the 3D variants (fixed): effective batch size differed from the configured one
• Frozen ResNet backbone's batch-norm statistics drifted during fine-tuning (fixed)
• Final-model validation was partly in-sample (fixed); E5 members s43-s46 were trained after all fixes
• Leaky leave-one-fold-out ensemble: first real run produced S1 AUC 0.956, above the honest band for subject-wise-split OASIS-2 studies (66-90% accuracy); the guardrail caught it before any number was recorded, and the design was replaced with same-fold ensembling (commit a4d75f2)

## Slide 10
Limitations and future work
Limitations
• 22 held-out visits: the accuracy CI covers a 40.9-point range
• Sensitivity 0.545 at the default threshold: 45.5% of demented visits are missed
• Single cohort and scanner family; no external validation yet
• Fold-to-fold std ±0.096 AUC: a retrain can move numbers by points
• Binary labels only (no CN/MCI/AD staging)
Future work
• Pre-training on OASIS-3 before fine-tuning on OASIS-2
• External validation on ADNI/AIBL (docs/plan/07)
• Conformal sets revisited at a calibration size that can support the coverage target
• Calibration beyond a single temperature
• Multi-class staging (CN/MCI/AD)

## Slide 11
Takeaways
• Honest longitudinal-MRI performance on OASIS-2 sits near AUC 0.74-0.78; published 95%+ claims usually contain same-patient leakage (Diagnostics 2025)
• Subject-level splits, CV as primary evidence, and a single declared test touch keep the evaluation honest
• Ensembling bought calibration (ECE 0.304 to 0.227), not accuracy (McNemar p = 1.0)
• Conformal coverage 54.5% vs the 90% target: failed, reported, not hidden
• Next: OASIS-3 pre-training and ADNI external validation

## Slide 12
Thank you
Data: OASIS-2, Marcus et al., Journal of Cognitive Neuroscience 22(12), 2010 (oasis-brains.org)
Method context: deep ensembles (NeurIPS 2017), leakage review (Diagnostics 2025), conformal prediction in clinical AI (CHEST 2025)
Research use only — not a medical device. Not a diagnosis.
NeuroSight AI · neurosight-e5 · report 2026-10-21